Abstract

Objective: To evaluate the transcutaneous bilirubin (by Minolta JM-103 Jaundice Meter) as a screening tool in near term (≥35 weeks) to term neonates with jaundice who presented to our hospital.

Methods: This prospective study which was done in Karbala teaching hospital for children, during the period from  April 1, 2010 through December 31, 2010, studied the correlation between the total serum bilirubin (TSB) measurements by (APEL BR-501 BILIRUBIN METRE) and the paired transcutaneous bilirubin (TcB) measurements (by Minolta JM-103 Jaundice Meter).

Result: Transcutaneous bilirubin showed a good correlation with TSB (r=0.827) in 107 neonates aged between 1 to 13 days old. The mean difference between TcB and TSB was -1.15.

Conclusion: Transcutaneous bilirubin (measured by JM-103) cannot be considered a substitute for laboratory measurement of TSB, but appears to be useful as a screening tool.

تقييم جهاز فحص البليروبين عن طريق الجلد

الخلاصة

اهداف البحث: لمعرفة كفاءه عمل جهاز فحص البليروبين عن طريق الجلد وامكان استخدامه في متابعه مرضى اليرقان الولادي.

الطرق: لقد اجريت هذه الدراسة المستقبلية لمئة وسبعه طفل من حديثي الولادة الذين لديهم يرقان ولادي في مستشفى الاطفال التعليمي في كربلاء للفترة من بداية نيسان 2010  الى نهاية كانون الاول 2010  وتمت اخذ القراءات بواسطه جهاز فحص البليروبين عن طريق الجلد ومقارنتها بنتائج الاجهزه المستخدمه في المختبرات الكيميائيه في المستشفى.

النتائج: لقد وجد ان خمسة وستون(57%) طفلا كانوا من الذكور مقابل( 43%) من الاناث وان نسبة الذكور للاناث كانت (1.32:1) وان معدل عمر الاطفال كانت (5.8) ايام مع مدى من( 1-13)يوم. كان معدل الوزن عند الولادة كان (3.036) كغم بمدى من( 2.00- 4.200 ) كغم وان (71%) ولدوا ولادة طبيعية بينما (29%) كانت ولادتهم قيصرية. لقد كانت رضاعة (69%) منهم طبيعية وكانت معدل النسبه بين نتائج الفحص عن طريق الجلد لنتائج المختبرات الكيميائيه هي(0.827).

الاستنتاج: نسبه المقارنه جيده بين نتائج جهاز فحص البليروبين عن طريق الجلد والاجهزه المستخدمه في المختبرات الكيميائيه في المستشفى لذلك من الممكن استخدام جهاز  فحص البليروبين عن طريق الجلد لمتابعه  الاطفال حديثي الولاده المصابين باليرقان.

ــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــــ

Introduction

N

oninvasive techniques for transcutaneous measurement of bilirubin (TcB) that are linearly correlated with serum levels may be used to screen infants [1,2], but determination of serum bilirubin level is indicated in patients with elevated age-specific transcutaneous bilirubin measurement, progressing jaundice, or risk for either hemolysis or sepsis [3].

Transcutaneous bilirubin (TcB) measurement devices use multiwavelength spectral reflectance from the skin surface and can be used to estimate total serum/plasma bilirubin (TB), and, thus, avoid blood sampling. However, TcB measurements are not reliable in infants undergoing phototherapy [2]. In addition, controversy still exists as to whether TcB testing is affected by skin pigmentation [4,5]. As a result, TcB measurements may be limited in these neonates (ie, infants treated with phototherapy, or darkly skinned), as the magnitude of hyperbilirubinemia cannot be accurately ascertained.

In several reports of racially and ethnically diverse groups of term and late preterm newborns, a close correlation between TcB and TB measurements was demonstrated [4]. However, at high levels of TB (˃15 mg/dL [257 micromol/L]), TcB measurements underestimate TB and need to be confirmed by standard laboratory methods [6,7]. As a result, TcB can replace TB in most circumstances when TB is <15 mg/dL (257 micromol/L) [2,6].

There are significant variations among different instruments for TcB measurement. When TcB is used clinically as a substitute for TB, values of new instruments should always be compared to TB performed by the laboratory to ensure good correlation [8].

The use of TcB screening reduces the number of blood tests for bilirubin determination compared to visual assessment without compromising detection of infants with significant TB values (eg, >75th percentile) [9,10].

The aim of this work is to evaluate a non invasive new transcautaneous jaundice meter  JM-103 as screening test and decrease unnecessary blood sampling in follow up of jaundice in neonatal period.

Patient and Methods

A prospective study was performed in Karbala teaching hospital for children. Term and near-term neonates presenting for  jaundice during the period from  April 1, 2010 through December 31, 2010 were included in this study with verbal consent from their parents.

Transcutaneous bilirubin (TcB) and capillary bilirubin assays were performed just once for each patient, with a maximum interval of 30 min between them. The age in days was recorded, along with gestational age, body weight, type of feeding, type of labor and gender. All infants weighed more than 2,000 g, had a gestational age of more than or equal to 35weeks, and were clinically healthy and  jaundiced. Infants who were receiving phototherapy or who had exchange transfusion were excluded.

All TcB measurements were performed by one investigator, using the bilirubinometer (jaundice meter Minolta JM-103, Drager). The measurements were obtained from the foreheads of the infants, while lying in a supine position. Gentle pressure was applied, while the fiber optic probe was placed against the forehead to exert even contact by the probe with the forehead skin.

The TSB assay was performed in the hospital chemical laboratory, using a direct spectrometer (APEL BR-501 BILIRUBIN METRE).              Both devices were calibrated before use daily, in accordance with the manufacturer’s instructions

Demographic data, TcB and TSB values were analyzed using SPSS v12. 0. The correlation coefficient between TcB and TSB was performed using Pearson Linear Regression Analysis. Error distribution, mean difference, and 95% confidence interval of the difference were evaluated.

Results

During the study period, a total of 107 TcB and TSB measurements were taken from term and near term neonates.

Their body weight ranged from 2.00 to 4.2 kg (mean=X=3.03, SD=0.50), gestational age  ranged from 35 to 41weeks (X =38.18, SD1.4), the age (figure1) ranged from 1 to 13 days (X=5.8 SD=2.6).

TcB ranged from 5.9 to 18.8 (X= 11.45 SD=2.8), TSB ranged from 6 to 19.7 (X=12.6 SD=3.37).

Figure 1 Age distributions

Male to female ratio 1.32:1 (Males 61, female 46). Feeding is by breast in majority (74, bottle feeding 29, mixed feeding is only 4). Seventy six has been delivered by normal vaginal route (cesarean section 46).

Table 1 demonstrates demographic character for 107 neonates.

Table 1 Demographics characters

|  | Minimum | Maximum | Mean | Std. Deviation |
| --- | --- | --- | --- | --- |
| Gestational age | 35.00 | 41.00 | 38.1869 | 1.44804 |
| B W in Kg | 2.00 | 4.20 | 3.0360 | .50159 |
| Age in days | 1.00 | 13.00 | 5.8037 | 2.60058 |
| TcB in mg/dl | 5.90 | 18.00 | 11.4561 | 2.88098 |
| TSB in mg/dl | 6.00 | 19.70 | 12.6065 | 3.37103 |

The Pearson correlation between TcB using JM-103 (TcB-JM) and TSB was significant (r = 0.827, p < 0.01). The mean difference between TSB and TcB-JM was -1.15 mg/dl [SD = 1.8 mg/dl, and 95% confidence interval of the mean (CI) -1.5 and -0.7].

The linear regression and curve estimations are seen in figures 2 and 3).

Figure 2 linear regression.

Figure 3 Linear curve.

Discussion

Visually determined jaundice may invite to blood sampling for serum bilirubin level. However, visual inspection, being subjective, usually inaccurate and unreliable and will result in a number of unnecessary blood sampling. Transcutaneous bilirubin device may provide an answer for an objective, noninvasive measurement, few skill-demanding, comfortable, fast and painless [10].

A new device, JM 103, introduced in 2003, provides the measurement result in a clinical unit and was reported to have less influence of skin maturation level and melanin than the previous JM 102 [11].

In our study the correlation coefficient between TcB (JM 103) and TSB is significant (r = 0.827, p < 0.01), in agreement with study done by maisels M J in China [12] (r=0.83), Tommy SK Lam [13] (r.83) and sanpavat [14] (r=0.80) but less so than that of Yasuda et al [15] (r = 0.93).

One factor which may account for the difference is different laboratory measurement of TSB due to inter-laboratory variability with a wide range of uncertainty is well documented [16].

According to our results, the mean difference between TSB and TcB-JM was-1.15 mg/dl [SD = 1.8 mg/dl, and 95% confidence interval of the mean (CI) -1.5 and -0.7].

The measurement  of TcB over forehead had a tendency to underestimate TSB level which in agreement with study of Snapvat [14].

The correlation coefficient does not provide information about clinical significance of diagnostic test, it was found that when TSB level increase the difference between values of TcB and TSB increase as well [17].  However, a lower level of TSB upon which treatment begins can cause frequent monitoring of blood sampling, a painful procedure with possible complications. TcB can be used as a screening test to determine the need for TSB measurement [18].

Although the accuracy of TSB is not clarified by comparison with high performance liquid chromatography (HPLC), which is considered to be the gold standard for bilirubin assay, the TSB, performed by clinical laboratory method, is the test a physician routinely relies upon to make decision in management of hyperbilirubinemia.

In the majority of studies evaluating TcB measurements, TSB was measured by laboratory instruments using diazo-based methods [19] that have interferences with hemoglobin and other intracellular compounds [20]. Because blood collected from newborns is often hemolysed, this could affect the accuracy of the clinical laboratory methods. Several studies have evaluated the accuracy and precision of TcB measurements compared to HPLC measurements; the HPLC measurement approach, unlike routine laboratory-based methods, is not subject to interference from hemolysis or lipemia[21.22].

These studies suggest that TcB measurements may be used not only as a screening device but also as a reliable substitute for clinical laboratory– based serum bilirubin measurements. This is in line with the recent National Academy of Clinical Biochemistry laboratory medicine practice guidelines, which concluded that the TcB meters currently available for clinical use in the US (BiliChek and JM-103) provide results comparable to laboratory TSB [23,24].

When comparing TcB and TSB measurements, it is also important to remember that the 2 methods of measurement may be evaluating different physiologic entities. Rubaltelli et al. [25] suggested that TcB methods measure the amount of bilirubin that has moved from the serum into the tissue, possibly mimicking the movement of bilirubin across the blood– brain barrier and into brain issue, whereas laboratory-based methods measure only bilirubin that is circulating in the blood. Thus, TcB may actually offer additional information not provided by TSB measurements, although this hypothesis remains to be proven.

TSB should be measured if the management plan would be altered by considering the TSB to be equal to TcB + 3 mg/dL (51 micromol/L). Therapy should be initiated while awaiting confirmatory results [26].

This study has few limitations:

The TSB was measured by the electro photometric device in our local chemical laboratory rather than the gold standard HPLC

We didn't measure TcB in preterm ˂ 35 weeks gestation, and we used only the forehead skin for measurement.

Conclusion

Noninvasive transcutaneous bilirubin assessment by JM 103 has demonstrated significant correlation when compared to total serum bilirubin measured by chemical laboratory method. It can be favorably used as a screening test to identify the need for serum bilirubin measurement.

References

1. Fay DL, Schellhase KG, Suresh GK: Bilirubin screening for normal newborns: a critique of the hour-specific bilirubin nomogram.  Pediatrics  2009; 124:1203-1205.

2. American Academy of Pediatrics, Subcommittee on Hyperbilirubinemia: Management of hyperbilirubinemia in the newborn infant 35 or more weeks of gestation, Pediatrics 2004, 114:297–316.

3. Maisels MJ, Kring E: Transcutaneous bilirubin levels in the first 96 hours in a normal newborn population of > or = 35 weeks’ gestation.  Pediatrics  2006; 117:1169-1173.

4. Wainer S, Rabi Y, Parmar SM, Allegro D, Lyon M. Impact of skin tone on the performance of a transcutaneous jaundice meter. Acta Paediatr. 2009;98(12):1909.

5. Bhutani VK, Gourley GR, Adler S, Kreamer B, Dalin C, Johnson LH. Noninvasive measurement of total serum bilirubin in a multiracial predischarge newborn population to assess the risk of severe hyperbilirubinemia. Pediatrics. 2000;106(2):E17.

6. Grohmann K, Roser M, Rolinski B, Kadow I, Müller C, Goerlach-Graw A, Nauck M, Küster H. Bilirubin measurement for neonates: comparison of 9 frequently used methods. Pediatrics. 2006;117(4):1174.

7. Engle WD, Jackson GL, Stehel EK, Sendelbach DM, Manning MD. Evaluation of a transcutaneous jaundice meter following hospital discharge in term and near-term neonates. J Perinatol. 2005;25(7):486.

8. Maisels MJ, Bhutani VK, Bogen D, et al. Hyperbilirubinemia in the newborn infant > or =35 weeks' gestation: an update with clarifications. Pediatrics 2009; 124:1193.

9. Kaplan M, Shchors I, Algur N, Bromiker R, Schimmel MS, Hammerman C. Visual screening versus transcutaneous bilirubinometry for predischarge jaundice assessment. Acta Paediatr. 2008;97(6):759.

10. Mishra S, Chawla D, Agarwal R, Deorari AK, Paul VK, Bhutani VK. Transcutaneous bilirubinometry reduces the need for blood sampling in neonates with visible jaundice.  Acta Paediatr. 2009;98(12):1916.

11. LY SIU, LW SIU, NS KWONG. Minolta JM-103 Jaundice Meter: A Screening Tool for Neonatal Jaundice in Chinese Neonates in Maternal and Child Health Centers. HK J Paediatr (new series) 2010; 15:204-213.

12. Maisels MJ, Ostrea EM Jr, Touch S, et al. Evaluation of a new transcutaneous bilirubinometer. Pediatrics 2004; 113:1628- 35.

13. Tommy SK Lam, KL Tsui, CW Kam. Evaluation of a point-of-care transcutaneous bilirubinometer in Chinese neonates at an accident and emergency department. Hong Kong Med J 2008; 14:356-60.

14. Sanpavat S, Nuchprayoon I. Transcutaneous bilirubin as a screening test to identify the need for serum bilirubin assessment. J Med Assoc Thai 2004; 87: 1193-8.

15. Yasuda S, Itah S, Isobe K, Yonetani M. Nakamura H, Nakamura M. et al. New transcutaneous jaundice device with two optical paths. J Perinat Med 2003; 31: 81-8.

16. Schreiner RL, Glick MR. Interlaboratory bilirubin variability. Pediatrics 1982; 69: 277-81.

17. Dalia Stonienė, Jūratė Buinauskienė, Eglė Markūnienė. The value of transcutaneous method of bilirubin measurement in newborn population with the risk of ABO hemolytic disease. Medicina (Kaunas) 2009; 45:792-796.

18. Briscoe L, Clark S, Yoxall CW. Can transcutaneous bilirubinometry reduce the need for blood tests in jaundiced full term babies? Arch Dis Child Fetal Neonatal Ed 2002; 86: F190-2.

19. Lo S, Doumas BT, Ashwood E. Performance of bilirubin determinations in US laboratories— revisited. Clin Chem 2004;50:190 – 4.

20. Kazmierczak SC, Robertson AF, Briley KP. Comparisonof hemolysis in blood samples collected using an automatic incision device and a manual lance. Arch Pediatr Adolesc Med 2002; 156:1072–4.  50 30.

21. Randeberg LL, Roll EB, Nilsen LTN, Christensen T, Svaasand LO. In vivo spectroscopy of jaundiced newborn skin reveals more than a bilirubin index. Acta Paediatr 2005; 94:65–71.

22. Kazmierczak SC, Robertson AF, Briley KP, Kreamer B, Gourley GR. Transcutaneous measurement of bilirubin in newborns: comparison with an automated Jendrassik-Grof procedure and HPLC. Clin Chem 2004; 50:433–5.

23. Kazmierczak S, Bhutani V, Gourley G, Kerr S, Lo S, Robertson A, Sena SF. Transcutaneous bilirubin testing. In: Laboratory medicine practice guideline: evidence-based practice for point-of-care testing. AACC Press 2007; 5–12.  (Accessed March 2009).

24. Nichols JH, Christenson RH, Clarke W, Gronowski A, Hammett-Stabler CA, Jacobs E, et al. Executive summary. The National Academy of Clinical Biochemistry Laboratory Medicine Practice Guideline: evidence-based practice for point-of-care testing. Clin Chem Acta 2007; 379:14–28.

25. Rubaltelli FF, Gourley GR, Loskamp N, Modi N,Roth-Kleiner AM, Sender A, Vert P. Transcutaneous bilirubin measurement: a multicenter evaluationof a new device. Pediatrics 2001; 107:1264–71.

26. Maisels, MJ. Use TcB as a screening tool for jaundiced newborns. AAP News 2004; 25:9.

# Text boxes

Transcutaneous Bilirubin Testing: A Screening Tool for Jaundiced Term and Near- Term Neonates

Zuhair Omran Easa

College of Medicine, Karbala University, Karbala, Iraq.

M J B

1.00

2.00

3.00

4.00

5.00

6.00

7.00

8.00

9.00

10.00

11.00

12.00

13.00

Age in days

0

5

10

15

20

Count