Processing of bioceramics

Technical ceramics are composed of raw materials generally as powder and of

natural or synthetic chemical additives, favoring either compaction (hot, cold or

isostatic), or setting (hydraulic or chemical) or accelerating sintering processes.

According to the formulation of the bioceramic and the shaping process used, we

can obtain ceramics, dense or with variable porosity, cements, ceramic depositions

or ceramic composites.

12.4.1. Massive ceramics

Molding, traditional or under pressure, and pressing, unidirectional or isostatic,

is part of the ceramist’s range of tools for obtaining the final form of the piece to be

made.

The heads of hip prothesis in zirconia or alumina are formed by isostatic pressing

(approximately 1,200 bars) from micronized powders. The sintering operation which

follows is carefully defined by heating and cooling parameters (temperatures and

rates), so as to obtain a finished product with a nil total porosity, especially in the

case of prostheses heads [MOC 93]. The choice of a micronized and reactive powder

514 Ceramic Materials

combined with isostatic pressing leads to this result [SED 93]. More recent

manufacturing methods borrowed from the electronic industry like strip molding can

also be used for making bone substitutes [RIC 98]. Pieces obtained are of small

thicknesses (from 100 µm to 2 mm) and of large surface (many tens of cm2).

Porosity is an important parameter for cell rehabilitation. It can be obtained by

different methods. The utilization of a pore forming agent is the method most widely

used. The pores generated by the release of a gaseous compound (H2O2, CO2),

combined either with slurry or a powder, are generally irregular and we prefer

calibrated organic substances (naphthalene, polystyrene, polyethylene), added to the

powder, which decompose slowly during heating. These techniques can be

employed for different types of ceramics. The use of calibrated polymer foams,

impregnated with slurry, has also been proposed and achieved successfully.

In parallel, more specific techniques have been sought for calcium phosphates

using smaller temperatures. Thus, porous ceramics in calcium phosphates

resembling coral skeletons can be obtained by the “replamineform” technique

(Interpore) in aqueous media [RIC 98]. Moreover, a hydrothermal process helping to

obtain apatitic ceramics at temperatures clearly lower than those used by ceramists

has been perfected and developed: this is the way of hydrothermal cements [DON

98]. Other methods using hydraulic cements have also been proposed [DRI 95].

These methods allow easy preparation, molding and machining of the sintered piece,

as well as a more flexible check of the composition and the additives. All these

methods lead, however, to porous compounds with weak mechanical performance.

Another way of material processing is developing; it is based on the biomimitic

processes aiming at imitating natural and biological processes and offers the

possibility of making these bioceramics at a temperature lower (ambient

temperature) than that of conventional or hydrothermal processes [GRO 96]. The

prospect of using these very low forming temperatures opens up possibilities for

mineral organic combinations with improved biological properties by the addition of

proteins and biologically active molecules (growth factors, antibiotics, anti-tumor

agents, etc.). However, these materials have poor mechanical properties which can

be improved, partially, by combinations with bonding proteins.