Zirconia

Zirconia doped with yttrium oxide has also been proposed as a substitute for

alumina in the heads of osteoarticular prostheses. Important developmental studies

have been made [CAL 95] and these heads have been commercialized. The main

advantages of zirconia compared to alumina are a greater failure strength, mainly

bending strength, as well as a good resistance to fatigue [DRO 97]. These properties

make the use of prostheses heads of very small dimensions possible, thus reducing

the wear debris. Besides, zirconia has a better coefficient of friction and a better

wear resistance, despite the quite controversial results. One of the unknowns is the

role of zirconia in the nucleation of calcium phosphates from supersaturated body

fluids. Moreover, as we have mentioned, alumina has the advantage of a longer use

and currently gives satisfactory results.

12.2.2.4. Vitreous carbon and diamond carbon

Vitreous carbon has several interesting physicochemical and biological

properties: it is light, resistant to wear and haemocompatible. It is essentially used

for making cardiac valves and replaces natural valves taken from pigs which have a

tendency to calcify and have a more limited life. Essential problems which remain

Bioceramics 497

are the formation of thromboses and bleeding due to the degradation of the junction

between the prosthesis and the artery. Mechanical valves appear, however, to be

tolerated in the long term [PET 99].

Diamond is an interesting coating from a biological point of view: it does not

induce any cytotoxic or haemolytic effects and can be used for various vascular

applications or in cardiac surgery [DIO 92].

12.2.2.5. Other ceramics

A number of other ceramics have been subjected to biomedical tests for

implantation, without currently being developed industrially. Among these ceramics,

we can cite silicon carbide, titanium nitrides and carbides, and boron nitride. TiN

has been suggested as the friction surface in hip prostheses. While cell culture tests

show a good biocompatibility, the analysis of explants shows significant wear,

related to a delaminating of the TiN layer [HAR 97]. Silicon carbide is another

modern day ceramic which seems to provide good biocompatibility and can be used

as bone implant [SAN 98].

12.2.3. Ceramics for specific uses

Ceramics for specific uses, in addition to their traditional properties, have

biological activity. We thus refer to bioactive ceramics. These ceramics are

essentially used as bone replacements and today form part of the daily practice of

orthopedic, maxillofacial and plastic surgeons. These are used in case a loss of bone

substance (tumor, important trauma, infection, etc.) requires to be filled. According

to the type and shape of the defect to be filled, the location of the implant and

mechanical stresses, different types of bioceramics with varying biological

properties are available. Biodegradable ceramics will be resorbed and replaced by

reconstructed tissue, whereas non-biodegradable ceramics are intended for a

permanent implantation. Very often, synthetic ceramics compete with natural

materials and it seemed more interesting to start this section with a description of the