2.3 Nd :YAG

The Nd :YAG laser is by far the most commonly used type of solid-state laser.

Neodymium-doped yttrium aluminum garnet (Nd :YAG) possesses a combination

of properties uniquely favorable for laser operation. The YAG host is hard, of

good optical quality, and has a high thermal conductivity. Furthermore, the cubic

structure of YAG favors a narrow fluorescent linewidth, which results in high

gain and low threshold for laser operation. In Nd :YAG, trivalent neodymium

substitutes for trivalent yttrium, so charge compensation is not required.

Physical Properties

Pure Y3Al5O12 is a colorless, optically isotropic crystal that possesses a cubic

structure characteristic of garnets. In Nd :YAG about 1% of Y3+ is substituted by

Nd3+. The radii of the two rare earth ions differ by about 3%. Therefore, with

the addition of large amounts of neodymium, strained crystals are obtained—

indicating that either the solubility limit of neodymium is exceeded or that the

TABLE 2.2. Physical and optical properties of Nd :YAG.

lattice of YAG is seriously distorted by the inclusion of neodymium. Some of the

important physical properties of YAG are listed in Table 2.2 together with optical

and laser parameters.

Laser Properties

The Nd :YAG laser is a four-level system as depicted by a simplified energy level

diagram in Fig. 2.2. The laser transition, having a wavelength of 1064.1 nm, originates

from the R2 component of the 4F3/2 level and terminates at the Y3 component

of the 4I11/2 level. At room temperature only 40% of the 4F3/2 population

is at level R2; the remaining 60% are at the lower sublevel R1 according

to Boltzmann’s law. Lasing takes place only by R2 ions whereby the R2 level

population is replenished from R1 by thermal transitions. The ground level of

Nd :YAG is the 4I9/2 level. There are a number of relatively broad energy levels,

which together may be viewed as comprising pump level 3. Of the main pump

bands shown, the 0.81 and 0.75μm bands are the strongest. The terminal laser

level is 2111 cm−1 above the ground state, and thus the population is a factor of

exp(_E/kT) ≈ exp(−10) of the ground-state density. Since the terminal level is

not populated thermally, the threshold condition is easy to obtain.

The upper laser level, 4F3/2, has a fluorescence efficiency greater than 99.5%

and a fluorescence lifetime of 230μs. The branching ratio of emission from 4F3/2

is as follows: 4F3/2 → 4I9/2 = 0.25, 4F3/2 → 4I11/2 = 0.60, 4F3/2 → 4I13/2 =

0.14, and 4F3/2 → 4I15/2 < 0.01. This means that almost all the ions transferred

from the ground level to the pump bands end up at the upper laser level, and 60%

of the ions at the upper laser level cause fluorescence output at the 4I11/2 manifold.

FIGURE 2.2. Energy level diagram of Nd :YAG.

At room temperature the main 1.06μm line in Nd :YAG is homogeneously

broadened by thermally activated lattice vibrations. The spectroscopic cross section

for the individual transition between Stark sublevels has been measured to

be σ(R2 − Y3) = 6.5 × 10−19 cm2. At a temperature of 295 K, the Maxwell–

Boltzmann fraction in the upper Stark sublevel is 0.427, implying an effective

cross section for Nd :YAG of σ

∗

(4F3/2 − 4I11/2) = 2.8 × 10−19 cm2. The effective

stimulated-emission cross section is the spectroscopic cross section times the

occupancy of the upper laser level relative to the entire 4F3/2 manifold population.

Figure 2.3 shows the fluorescence spectrum of Nd3+ in YAG near the region of

the laser output with the corresponding energy levels for the various transitions.

The absorption of Nd :YAG in the range 0.3 to 0.9μm is given in Fig. 2.4(a). In

Fig. 2.4(b) the absorption spectrum is expanded around the wavelength of 808 nm,

which is important for laser diode pumping. Thermal properties of Nd :YAG

are summarized in Table 2.3. Under normal operating conditions the Nd :YAG

laser oscillates at room temperature on the strongest 4F3/2 → 4I11/2 transition

at 1.0641μm. It is possible, however, to obtain oscillation at other wavelengths

by inserting etalons or dispersive prisms in the resonator, by utilizing a specially

FIGURE 2.3. Fluorescence spectrum of Nd3+ in YAG at 300K in the region of 1.06 μm

[10], [11].

designed resonant reflector as an output mirror, or by employing highly selective

dielectrically coated mirrors. These elements suppress laser oscillation at

the 1.06μm wavelength and provide optimum conditions at the wavelength desired.

With this technique laser systems have been built which utilize the 946 and

1330 nm transitions.

FIGURE 2.4. (a) Absorption spectrum of Nd :YAG from 0.3 to 0.9μm and (b) expanded

scale around 808 nm.

TABLE 2.3. Thermal properties of Nd :YAG.

FIGURE 2.8. Output from a Nd:YVO4 and Nd :YAG laser as a function of diode pump

temperature and wavelength [13].

3.6.2 Diode Side-Pumped Nd : YAG Laser

The system illustrated here produces an energy per pulse of about 0.5 J at a repetition

rate of 40 Hz. Critical design issues for this laser include heat removal

from the diode arrays and laser rod and the overlapping of the pump and resonator

mode volumes. In side-pumped configurations, laser-diode arrays are not

required to be coherent, and pump power can be easily scaled with multiple arrays

around the circumference of the rod or along its axis. Instead of one diode array

pumping the laser crystal, this particular laser employs 16 diode arrays located

symmetrically around the rod. As shown in Fig. 3.19 the diode pumps are arranged

in four rings, each consisting of four arrays. Since each array is 1 cm long,

the total pumped length of the 6.6 cm × 0.63 cm Nd :YAG crystal is 4 cm. This

arrangement permits the incorporation of large water-cooled heat sinks required

for heat dissipation, and it also provides for a very symmetrical pump profile. An

eight-fold symmetry is produced by rotating adjacent rings of diodes by 45◦. A

photograph of the extremely compact design is also shown in Fig. 3.19. The symmetrical

arrangement of the pump sources around the rod produces a very uniform

pump distribution, as illustrated in Fig. 3.20. The intensity profile displays

the fluorescence output of the rod taken with a CCD camera. In Fig. 3.21, the

output versus optical pump input is plotted for long pulse multimode and TEM00

FIGURE 3.19. Cross section (left) and photograph (right) of diode-pumped Nd :YAG laser

head.

FIGURE 3.20. Pump distribution of a 16-diode array side-pumped Nd :YAG crystal (each

line represents a 10% change in intensity).

mode operation. Also shown is the output for Q-switch TEM00 mode operation.

The resonator configuration for the long pulse, multimode operation is depicted

in Fig. 3.22. The TEM00 mode performance was achieved with a variable reflectivity

mirror and a concave–convex resonator structure which will be described in

Chapter 5.

The multimode laser output can be expressed by

FIGURE 3.21. Output versus input energy for diode-pumped Nd :YAG oscillator.

FIGURE 3.22. Resonator configuration for multimode operation.

where Eopt is the optical pump energy from the 16 diode arrays. The electrical

input energy Ein required to achieve Eopt is

Combining these two output–input curves relates the laser output with the electrical

input energy

The slope efficiency of the laser is 25%, and the overall electrical efficiency at

the maximum output of 460 mJ per pulse is 16%. The optimum output coupling

was experimentally determined. Figure 3.23 shows a plot of the laser output for

different values of the reflectivity. The different efficiency factors of the system

FIGURE 3.23. Oscillator output vs. mirror reflectivity (calculated curve based on Topt =

0.2, g0l = 1.4, δ = 0.03)

TABLE 3.1. Energy transfer efficiencies of a large

diode-pumped Nd :YAG oscillator.

are listed in Table 3.1. The slope efficiency of the diode array was measured with

a power meter. The absorbed pump radiation in the 6.3mm diameter crystal was

calculated from a computer code.

The coupling efficiency ηC follows from Topt = 0.20 and the measured roundtrip

loss of δ = 2.2%. The gain/mode overlap efficiency ηB was estimated by

comparing the beam profile with the pump distribution in the laser rod. The final

value was adjusted for the product of the η terms to agree with the measured slope

efficiency of 25%. For TEM00 mode operation, the major difference is a substantially

reduced value for ηB as a result of a smaller beam. All other parameters

remain unchanged.

3.6.3 End-Pumped Systems

In this so-called end-pumped configuration, the radiation from a single laser diode

or diode bar is focused to a small spot on the end of a laser rod. With a suitable

choice of focusing optics, the diode pump radiation can be adjusted to coincide

with the diameter of the TEM00 resonator mode. The end pumping configuration

thus allows the maximum use of the energy from the laser diodes.

Using this longitudinal pumping scheme, the fraction of the active laser volume

excited by the diode laser can be matched quite well to the TEM00 lasing volume.

A solid-state laser pumped in this manner operates naturally in the fundamental

spatial mode without intracavity apertures.

Many excellent systems with extremely compact packaging and high efficiency

that utilize the mode matching pump profile are commercially available.

By way of illustrating the concept, we will consider the configuration illustrated

in Fig. 3.24 which was originally proposed by Sipes [5]. The output from a

diode array with 200mW output at 808 nm is collimated and focused into a

1 cm long × 0.5 cm diameter 1% Nd :YAG sample. The resonator configuration

is plano-concave, with the pumped end of the Nd :YAG rod being coated for

high reflection at 1.06μm, and with an output coupler having a 5 cm radius of

curvature and a reflectivity at 1.06μm of 95%.

FIGURE 3.24. End-pumped laser oscillator.

Figure 3.25 displays electrical input power versus 1.06μm output power for

the configuration illustrated in Fig. 3.24. We see that for approximately 1W of

electrical input power, 80mW of Nd :YAG output is measured.

The electrical slope efficiency of the laser is 13%, and the overall efficiency at

80mW output is 8%. Also shown in Fig. 3.25 is the performance of the diode array

that has a slope efficiency of 34% and an overall efficiency of 22% at 220mW

output. The energy transfer steps of the laser are listed in Table 3.2. The slope efficiency

ηP of the diode laser output is a measured quantity. The transfer efficiency

ηT includes the collection of diode radiation by the lens system and reflection

losses at the surfaces. The pump radiation is completely absorbed in the 1-cm

long crystal, i.e., ηa = 1. The coupling efficiency ηC follows from the measured

FIGURE 3.25. Nd :YAG laser and diode output versus electrical input power.

TABLE 3.2. Energy transfer efficiencies of

an end-pumped Nd :YAG oscillator.

1% one-way loss and an output coupling of 5%. As stated by the authors, the

large ellipticity of the diode array beams made it difficult to focus the entire pump

beam into the laser resonator mode. The value of ηB is therefore lower than can

be achieved with an optimized system. The measured slope efficiency of the laser

is the product of the η-terms listed in Table 3.2.

This example of an end-pumped laser demonstrates the high overall performance

at TEM00 mode which can be achieved with this type of system. The attractive

features of end-pumped lasers are a very compact design combined with

high beam quality and efficiency, as a result of the good overlap between the

pumped region and the TEM00 laser mode.

Figure 3.26 displays a photograph of a typical end-pumped laser. The optical

components shown separately are the laser diode array, two cylindrical and two

FIGURE 3.26. Photograph of (a) assembled end-pumped laser and (b) disassembled optical

components. [Fibertek, Inc.].

FIGURE 3.27. Laser output as a function of pump energy at the laser crystal.

spherical lenses, an Nd :YAG crystal, polarizer, Q-switch, and output coupler.

The end mirror of the resonator is coated onto the laser crystal. The diode array

comprises three quasi-cw 1 cm bars. The three-bar stack had a maximum output

energy of 27 mJ in a 200μs pulse. At that pulse energy, the laser output was 6.7 mJ

in the TEM00 mode and 8.4 mJ multimode (Fig. 3.27). The pulse repetition was

50 pps. The resonator loss was determined by measuring the pump threshold as a

function of the reflectivities of the output coupler, as described in Section 3.4.2.

The measurement yielded a value of 2.4% for the resonator round-trip loss. The

maximum output was obtained for an output coupler reflectivity of 95%. The

optical design of this laser is depicted in Figs. 6.39.

Summary

Basically a laser oscillator consists of a medium with optical gain inside a resonator

formed by two mirrors. The two mirrors perform the function of optical

feedback by reflecting the optical radiation back and forth through the gain

medium.

The oscillator starts to lase after having been triggered by some spontaneous

radiation emitted along the axis of the resonator. This radiation is amplified and a

fraction leaks out of the output mirror. Most of the radiation is reflected and amplified

again as it passes through the active material. To continue oscillating, the

gain per double pass must exceed the losses. As the intensity builds up inside the

resonator, the gain is reduced because of gain saturation. A stable operating point

is reached when the round trip gain and the losses are equal. Dynamic interactions

between the gain of the active medium and the photon flux in the resonator causes

damped oscillations around the steady state value.

In order to achieve maximum output from an oscillator under given operating

conditions the reflectivity of the output mirror has to be optimized. Increasing the

output coupling (T = 1 − R) is equivalent to decreasing the feedback factor and

increasing the loss in the resonator. The limiting condition is that feedback must

be sufficiently large to compensate for the internal and output coupling losses in

the system. The transmission of the output mirror will have an optimum value that

maximizes the output power from the oscillator. This value is between zero and a

transmission at which the losses exceed the gain and the oscillator will cease to

lase.

Threshold and laser output as a function of materials and systems parameters

are discussed in this chapter. The conversion of electrical input energy to laser

output proceeds through a number of steps that can be related to the pump source,

pump radiation transfer system, laser medium, and optical resonator. The efficiency

of the energy conversion or transfer process associated with each step

contributes to the overall efficiency of the laser. The energy flow within a laser

oscillator is analyzed in detail, and an illustration of the practical realization of

different oscillators concludes the chapter.

Solid-State Lasers:

A Graduate Text

Walter Koechner

Michael Bass

_c 2003 Springer-Verlag New York, Inc.

6.1.3 Laser Diodes

The high pump efficiency compared to flashlamps stems from the excellent spectral

match between the laser diode emission and the Nd absorption bands at

808 nm. Actually, flashlamps have a radiation efficiency (40–60%) comparable

to laser diodes (35–50%); however, only a small fraction of the radiation match

by the laser material absorption bands. In contrast, the output wavelength of laser

diodes can be chosen to fall completely within an absorption band of a particular

solid-state laser. This has also the added benefit of reduced thermal loading

of the laser crystal, which leads to better beam quality. The directionality of the

diode output and the small emitting area provides great flexibility of shaping and

transferring the pump radiation from the source to the laser medium that created

whole new design architectures, such as end-pumped systems, microchip lasers,

and fiber lasers. Diode laser pumps are also an enabling technology for new laser

materials. The most prominent laser materials that are pumped with diode pump

sources can also be pumped with flashlamps. However, a number of very useful

materials such as Nd :YVO4, Yb :YAG, and Tm:YAG have only reached prominence

as a result of diode laser pumps.

Internal Structure of Laser Diodes

The most basic form of a diode laser consists of a semiconductor in which one

part is doped with electron donors to form n-type material and the other part is

doped with electron acceptors (holes) to produce p-type material. Application of

a negative voltage to the n-material and a positive voltage to the p-material drives

electrons and holes into the junction region between the n- and p-doped material.

This is called forward biasing the pn-junction. In the junction electron–hole

recombination takes place. In this process, electrons from a higher-energy state

(conduction band) transfer to a lower-energy state (valence band) and this releases

energy in the form of photons. In order to produce stimulated emission in the junction

area, population inversion and optical feedback are required. Passing a high

current through the junction area provides a population inversion. A Fabry–Perot

resonator, formed by cleaving the chip along parallel crystal planes, provides optical

feedback. Gain in laser diodes is very high; therefore only a small length

of active material is needed. Typical edge emitting laser chips have active layers

about 500μm long. In order to increase efficiency and output power, actual laser

diodes contain a number of different layers with the active region sandwiched

in between. Semiconductor-laser technology has produced an amazing variety of

new device structures made possible by sophisticated growth techniques such as

MetallOrganic Chemical–Vapor Deposition (MOCVD) and molecular beam epitaxy

(MBE) [7]. Most techniques differ in the way confinement is achieved regarding

the width and depth of the gain region and mode volume of the optical

radiation.

Figure 6.9 shows the dimensions of a typical 1 cm monolithic linear array (top)

and a cross section of the device (bottom) that is made up of a stack of several

semiconductor layers with varying thickness and doping levels. The thickest

layer on the bottom is the n-GaAs substrate onto which the other layers are deposited.

The top layer is p-GaAs. Sandwiched between several other layers is

the very thin active region of the pn-junction. The top and bottom surfaces of the

whole stack are coated with metallic contacts for electrical connection to the drive

electronics. The active region is divided into many parallel and narrow stripes

by restricting the current in the lateral direction. The width of the active region

has to be restricted to avoid the onset of amplified spontaneous emission which

would propagate in off-axis directions and compete with the laser beam for the

available population inversion. In the gain guided design illustrated in Fig. 6.9(a)

particular emitter width is achieved by deposition of regions of high resistivity

material beneath the contact area. Therefore the current is channeled through narrow

low-resistance areas of the p-GaAs surface layer (white rectangular area in

Fig. 6.9(b)). Electrical current injected through the forward biased pn-junction

creates the population inversion and gain in the active region. The distribution of

the charge-carriers injected in the active region, and therefore the gain distribution,

is determined by the width of the current channel. The voltage across the

junction is on the order of 2V for GaAs/AlGaAs materials, and the slope efficiency

is about 1 W/A. Laser emission, which is between 790 and 860 nm for

FIGURE 6.9. Gain-guided, single quantum-well separate confinement heterostructure

stripe laser [8]. (a) Typical dimensions of a 1 cm bar and (b) detail of internal structure.

(Thickness of layers is greatly exaggerated).

GaA/AlGaAs laser diodes, is perpendicular to the planar front facet of the laser

diode bar shown in Fig. 6.9.

A 1 cm bar contains 10 to 50 independent laser emitters, all radiating in the

same direction perpendicular to the planar front of the structure. The width of

each emitter (white rectangle in Fig. 6.9(b)) is between 50 to 200 μm. The spacing

between adjacent emitters ranges from 200μm to 1mm. In so-called broad

area emitters each emitter consists of a single stripe. In other designs the total

width is subdivided in subarrays of 10 to 20 closely spaced narrow-stripe emitters

for better lateral mode control. The width of, and spacing between, individual

emitters is dictated by the average power, peak power, or duty cycle requirements

and the type of cooling applied to the bar.

Radiation in the vertical direction of the pn-junction is confined by a wave

guide structure comprised of several layers of differently doped materials. The

active layer shown in Fig. 6.9 is sandwiched between two pairs of layers having

a different concentration of aluminum. The main purpose of the innermost

pair of layers is to confine the carriers to the active region, whereas the purpose

of the outer layers is to confine the optical beam. In a conventional double

heterostructure pn-junction the active region has a thickness around 0.1–0.3μm.

The quantum-well active region illustrated in Fig. 6.9(b) is an order of magnitude

thinner. Extremely high gains on the order of 3 to 5×104 cm−1 and low thresholds

can be achieved in such a thin layer.

A quantum well is a thin layer of semiconductor located between two layers

with a larger bandgap. Electrons in the quantum-well layer lack the energy to

escape and cannot tunnel through the thicker surrounding layers. The quantumwell

layer in Fig. 6.9(b) has a composition of gallium and aluminum indicated

by x, which defines the emission wavelength. A higher aluminum concentration

increases the bandgap and shifts the output toward shorter wavelength. The

quantum-well structure is sandwiched between two thick layers of a composition

which contains a higher concentration of aluminum (y > x). The higher aluminum

concentration increases the bandgap, thereby defining the quantum well,

and the large thickness of the layer prevents tunneling of the carriers out of the

quantum well. Instead of a single quantum-well active layer, several quantum

wells can be sandwiched one above the other to generate multiple quantum-well

stacks.

The thin active region incorporating a quantum-well active layer structure provides

low threshold and high electrical-to-optical efficiency. However, such a very

small emitting surface poses one problem since the power from a diode laser is

limited by the peak flux at the output facet. One can increase the output from

these devices by spreading the beam over an area that is larger than the active

layer or gain region. The standard approach is to deposit layers next to the active

layer, each of which has a slightly lower refractive index than the active layer,

thus making a wave guide of the active layer. In the separate carrier and optical

confinement heterostructure (SCH) design shown in Fig. 6.9(b), the refractive index

boundary is abrupt, at the layer boundary. Alternatively, the refractive indices

of the surrounding layers may be graded, forming a Graded-Index and Separate

carrier and optical Confinement Heterstructure (GRINSCH).

In the single quantum-well structure depicted in Fig. 6.9(b), the effective aperture

of the laser output is 0.1 to 0.4μm vertical to the pn-junction, which is about

a factor 20 larger than the gain region. This substantially reduces the energy density

at the output facet and enhances reliability by minimizing catastrophic facet

damage.

The small white rectangle in Fig. 6.9(b) indicates the gain region, whereas the

large rectangle surrounding it represents the optical aperture of one emitter. The

drawing in Fig. 6.9 exaggerates the vertical dimension by several orders of magnitude.

The emitting aperture is, in reality, a wide and very narrow horizontal line

with a width up to 200μm and a thickness of no more than 0.4μm.

As described above, the beam cross section of a single emitter is given by the

width of the current channel in the plane of the pn-junction and by the waveguide

structure in the vertical direction. The front and rear facets, separated by about

500μm, define the length of the Fabry–Perot resonator. The facets, which provide

optical feedback and output coupling, are created by cleaving the bars from the

processed semiconductor wafer. Cleaving along directions dictated by the crystal

structure generates extremely planar and parallel surfaces. The facets are dielectrically

coated to provide a high reflectivity at the rear facet and an appropriate

output transmission at the front facet.

Diode Pump Source Configuration

Depending on the output-power requirement of the laser and the particular pump

configuration chosen, one can select a diode pump source from a number of standard

commercial designs. In accordance with the average power capability we can

classify diode pumps into four groups:

(a) Laser diodes with emitting apertures between 100 and 500μm and output

powers up to 5W.

(b) Single 1 cm long bars yielding a cw output up to 40W.

(c) Stacked 1 cm bars, forming two-dimensional arrays that are conductively

cooled and mounted to a common heat sink.

(d) Stacked bar two-dimensional arrays with individual, internal liquid cooling

for each bar.

Depending on the number of bars and the type of cooling employed, several kilowatts

of cw output and tens of kilowatts of peak power in quasi-cw mode can be

obtained from two-dimensional arrays.

For small solid-state lasers a typical pump source is a laser diode as illustrated

in Fig. 6.10. The diode has a total output of 2W in the cw mode. The output is

emitted from an area of 200×1.0μm with a beam divergence of 40◦ by 10◦. The

device can be coupled to an optical fiber or mounted on a thermoelectric cooler.

In the latter configuration shown in Fig. 6.10 the array is mounted on a pyrami-

FIGURE 6.10. Diode laser mounted on a thermoelectric cooler (SDL, Inc.).

dal heat sink that connects to a thermoelectric cooler. Heat dissipates through the

base, while the radiation emerges from the top. The thermoelectric cooler permits

temperature tuning of the output to match the absorption lines of the solid-state

laser. In a GaAlAs diode laser, the wavelength changes with temperature according

to 0.3 nm/◦ C. Combining several of the 200μm wide emitters up to 5W of

output can be achieved in similar packages.

Linear Arrays

The most common pump source is a 1 cm diode bar, as illustrated in Figs. 6.11

and 6.12. It consists of a monolithic chip of linear diode arrays that is bonded

to an alloy submount. An insulated wire bond plate is soldered to the submount

behind the laser diode array. Electrically the individual diodes in the array are

FIGURE 6.11. Basic structure of a 1 cm bar.

FIGURE 6.12. Monolithic 1 cm bar for cw operation [9].

all connected in parallel to a wire bond plate. The bonding of the bar to the heat

sink is made through the thinner p-doped side to minimize the thermal resistance

between the gain medium and the heat sink.

In the fabrication of laser bars 10 to 50 independent laser emitters are serially

repeated on a single substrate to form a monolithic bar structure. Practical processing

considerations limit the length of such bars to 1 cm. In the example shown

in Fig. 6.12, the device is composed of 20 emitters spaced on 500μm centers and

produces an output of 20W in the cw mode. Recent devices are capable of up to

40W output power in a 1 cm bar. In the particular design shown in Fig. 6.12 each

emitter is further subdivided into a number of narrow stripes. The spacing of the

emitters, which amounts to a fill factor of 20% for the device, is determined by

thermal limitations.

The use of diode lasers for pumping pulsed neodymium lasers requires a pulse

length on the order of several hundred microseconds. One limit of the output of

diode lasers is the temperature of the junction and associated thermal runaway

condition that raises the local temperature of the facet to the melting temperature.

Because of the small volume of material involved, this temperature rise occurs in

less than 1μs. Therefore, the power limit of diodes in the long-pulse mode is the

same as for cw operation. Thus, long-pulse operation (_ 1μs) is called quasi-cw

operation. The peak power obtained from a 1 cm bar can be increased over the cw

value by increasing the number of stripe lasers; this is achieved by increasing the

fill factor to close to 100%. Of course, the duty cycle has to be limited such that

the average power from the bar is about the same as for a cw bar. Operated in the

quasi-cw mode, a single bar typically generates 60W peak power, or 12 mJ in a

200μs long pulse, although devices up to 100Ware available at reduced lifetime.

Depending on cooling conditions, duty cycles on the order of 30% are possible.

Therefore, the maximum average output power for cw and quasi-cw devices is 20

to 40Wfor a 1 cm bar. Microlenses can be mounted in front of a bar to concentrate

pump radiation into a fiber. Typically a 20W cw bar can be coupled to a 400μm

fiber with 16W available at the output end of the fiber. The ability to mount the

pump source away from the laser onto a heat dissipating structure is very useful

in some applications.

Two-Dimensional Arrays

A number of 1 cm bars can be combined to form a two-dimensional array. The key

issues are the mounting technique and the efficient removal of waste heat from

the laser emitter. In the so-called rack-and-stack assembly technique, laser diode

bars are first mounted on individual submounts which are then stacked to produce

two-dimensional arrays. The spacing (pitch) between adjacent bars depends on

the average output power and on the cooling method.

For quasi-cw arrays one common design approach uses thin copper heat

sink spacers between the bars. The assembly of sandwiched bars and spacers

is mounted onto a common heat sink. Such a device is shown in Fig. 6.2(a). This

internal passively cooled diode stack is attached via the two mounting holes to

a larger cold plate or water-cooled heat sink. Electrically the bars are connected

in series, the two large copper leads extending from the back of the device in

Fig. 6.2(a) are for connection to a power supply. The thickness of the copper

heat spreader, and therefore the stacking density, depends on the duty cycle of

operation. Higher average power operation requires thicker heat sinks. For lowduty-

cycle operation around 3%, the bars can be densely packaged with only

0.4mm space in between; this allows a power density of 1.5 kW/cm2. At a 6%

duty cycle, the spacing increases to 0.8mm and 800 W/cm2 can be achieved.

Operation of a 10 stack array, with 200μs pulses at 100 Hz, produces typically a

peak power of 600W at 40% efficiency.

Multibar arrays for cw or high-duty cycle-pulsed operation require active cooling

close to the junction of the individual emitters. This can be achieved by using

microchannel cooling technology whereby each diode bar has its own liquid

heat exchange submount. In one approach individual bars are separated by

cooling plates apprxoimately 1mm thick, through which cooling water is channeled.

Using this approach, bars can be stacked with a 2mm bar-to-bar pitch.

Figures 6.2(b) and 6.13 depict internally liquid-cooled two-dimensional arrays.

Arrays comprised of 20 cw bars have achieved 1kW output in a 4 cm2 emitting

area. Pulsed arrays can be operated up to 30% duty cycle with internal active cooling.

The most powerful multibar pump modules generate peak powers in excess

of 30kW and cw powers of several kilowatts.

Spatial Profile

The output beam emerging from the facet of a diode laser is in the form of a

light cone that is oval in cross section, as illustrated in Fig. 6.14. In the vertical

direction with respect to the pn-junction the beam is diffraction limited and has

a divergence of about 40◦ FWHM. The Gaussian beam appears to emerge from

a 1μm aperture. If one introduces in (5.6) a wavelength of λ = 0.8μm and a

spot size for a Gaussian beam of 1μm the beam divergence at the 1/e2 points is

2θ = 57◦.

FIGURE 6.13. High-duty cycle internally liquid-cooled stacked bars (SDL, Inc.); (a) without

and (b) with microlens array.

FIGURE 6.14. Radiation pattern from a laser diode [10].

Collimation of such a highly divergent beam requires optics with a numerical

aperture of NA = sin(2θ/2) = 0.5 or higher. Optical systems with such a high

numerical aperture are said to be fast. From this it follows that the direction of the

output beam from a laser diode in the vertical direction of the emitting stripe is

termed the fast axis.

In the plane parallel to the pn-junction, the beam divergence is much smaller

and is about 10◦ FWHM. However considering the width of the emitter which is

between 50 and 200μm, this is many times diffraction limited. For example, a

Gaussian beam emerging from an aperture of 150μm would have a divergence

of 0.4◦ at the 1/e2 points. The output consists of many high-order spatial modes,

termed lateral modes in the literature. The number of lateral modes depends on

the design of the emitter (i.e., broad area or subdivided into narrow stripes), cavity

length, and various material parameters. Compared to the vertical direction, the

smaller divergence of the beam in the plane of the pn-junction allows the use of

optics with a smaller numerical aperture. Therefore the beam in this direction is

referred to as the slow axis.

Reformatting the highly astigmatic pump beam emitted from a line source into

a circular beam, which is required for end-pumped lasers, is a challenging task.

End-pumped lasers require focusing of the pump radiation to a small spot size for

mode matching or for coupling into an optical fiber. In Section 6.2.3 a number of

different optical schemes that accomplish this task will be discussed.

Diode Laser Lifetime

The long lifetime of laser diodes as compared to arc lamps is an important consideration

for pumping solid-state lasers with these devices. We will distinguish

between catastrophic failures and normal degradation during operation. Laser

diodes, like all semiconductors, are very susceptible to damage by electrostatic

discharges or high-voltage transients. Therefore handling of these components requires

electric discharge control such as wearing wrist straps and grounded work

surfaces. Also great care has to go into power supply and driver design to protect

this very expensive pump source from transients, current surges, and high voltage.

Operation at excessive currents will lead to catastrophic mirror facet damage.

Also, insufficient cooling combined with high currents can melt the solder joints

or vaporize the wire bonds within the diode structure.

Under proper operating conditions laser diodes degrade in a fairly predictable

manner which results in a decrease in output over time. The higher the operating

current and/or the operating temperature, the faster the degradation. The decrease

in output over time can be caused by structural defects, which spread through

the laser diode forming light absorbing clusters. Dark lines or spots in the output

beam are a manifestation of this damage. The degradation of AlGaAs laser diodes,

the most common pump source, is usually attributed to oxidation and migration

of aluminum under high-power operation. Structural defects can also be caused

by thermal stress due to different thermal expansion coefficients of semiconductor

and heat sink material.

Facet damage due to the environment or high-radiation flux is another source

of degradation, so is diffusion of solder into the semiconductor lattice over time

caused by the high-current density. After-screening of the laser diodes during a

burn-in period performed by the manufacturer has eliminated early failures due to

production problems, and very long operational lifetimes can be achieved. Lifetimes

of cw devices are typically on the order of 104 hr. In most commercial systems

requiring long life, diode arrays are typically operated at 75% of their rated

output. Quasi-cw bars and stacked arrays achieve lifetimes of several billion shots.

Spectral Properties

The spectral properties of laser-diode arrays that are most critical for the pumping

of solid-state lasers are the center wavelength and the spectral width of the

emission and the wavelength shift with temperature.

The wavelength of a laser diode is inversely proportional to the energy difference

_E between the conduction and valence band

where h is the Planck constant and c is the speed of light. The bandgap depends

on the crystalline structure and chemical composition of the semiconductor. A binary

compound such as GaAs has a fixed wavelength. Adding a third element, by

substituting a fraction of gallium with aluminum in GaAs, changes the bandgap.

In a compound semiconductor, such as GaAlAs, the wavelength of the emitted

light can be tailored over a wide range by changing the ratio of aluminum

and gallium. Adding aluminum increases the bandgap, producing lasers with

shorter wavelengths. Most commercial GaAlAs lasers have wavelengths of 750

to 850 nm. Neodymium ions have substantial absorption in the vicinity of 808 nm,

which is the emission wavelength of diode lasers with Ga0.91Al0.09As active regions.

As far as the pumping of neodymium lasers is concerned, the output wavelength

can be tailored to the peak absorption by adjustments of the aluminum

concentrations according to [11]

where x is the aluminum concentration.

Compositional changes and temperature gradients within an array lead to a

broader spectral output for the whole array as compared to a single device. A typical

spectral width of a single bar or small stacked array is on the order of 4 to

5 nm. State-of-the-art performance is about 2.2 nm for a 10 bar array. The bandwidth

of the Nd :YAG absorption line at 808 nm is 2 nm for an absorption coefficient

larger than 3.8 cm−1. Similarly, in Nd : YLF, the absorption line is about

2 nm wide for absorption coefficients larger than 2 cm−1. The excellent spectral

overlap between the absorption line of Nd :YAG at 808 nm and the emission of a

two-dimensional array of GaAlAs laser diodes is illustrated in Fig. 6.15.

FIGURE 6.15. Spectral overlap between Nd :YAG absorption and emission spectrum of a

10 bar array [12].

FIGURE 6.16. Absorption efficiency of Nd :YAG as a function of absorption length for a

pump wavelength of 808 nm [13].

In diode-pumped lasers that have only a short absorption path, it is important

to have a narrow spectral emission from the diode array in order to absorb most

of the pump radiation. In optically thick materials, such as large diode-pumped

lasers with rod or slab dimensions on the order of 10–15 mm, spectral width becomes

less important because eventually all the pump radiation gets absorbed. In

Fig. 6.16 the absorption of diode-pumped radiation for Nd :YAG is plotted versus

optical thickness with spectral width as a parameter.

For laser materials with an absorption band between 750 and 850 nm, pumping

with GaAlAs diodes is the obvious choice because of the availability and high

state of development of these devices. Lasers that can be pumped with GaAlAs

diodes include neodymium doped materials and Tm:YAG, as listed in Table 6.2.

For pumping Yb :YAG or yttrbium-sensitized lasers such as the Er : Yb : glass, a

TABLE 6.2. Important lasers and diode pump sources.

pump source around 0.98μm is needed. For this wavelength, so-called strainedlayer

superlattice lasers based upon InGaAs on GaAs substrates have been developed

with emission wavelengths between 0.95 and 0.98μm [14]. Traditional

diode lasers are restricted to materials that can be lattice-matched to GaAs or InP

substrates because unmatched materials are prone to defects and quickly degrade.

However, it was discovered that very thin layers—below a critical thickness of a

few tens of nanometers—can accommodate the strain of a small lattice mismatch.

This discovery led to the development of a single, thin strained layer used as a

quantum well which has an emission wavelength not achievable with GaAlAs or

InGaAsP laser diodes.

On the short-wavelength end of available diode pump sources are the visible

AlGaInP diode lasers operating in the 640–680nm region. These devices have

been used to pump chromium-doped lasers, such as Cr : LiCaAlF6 (LiCAF) and

Cr : LiSrAlF6 (LiSAF).

Input Versus Output Characteristics

The optical losses in the diode structure determines the threshold current is that

has to be exceeded before amplification can take place. A diode laser that is operating

above threshold will exhibit a linear relationship between output power and

electrical current as shown in Fig. 6.17. The optical output power as a function of

current input can be expressed by

where ηd is the differential quantum efficiency characterized by the number of

photons emitted per injected electrons, _E is the bandgap of the recombination

region, and e is the electron charge.

FIGURE 6.17. The cw output power versus input current for a 1 cm bar.

The slope efficiency σs, determined from the slope of the output power versus

input current, follows from (6.15a)

Figure 6.17 depicts the output power versus current for a 10Wcw bar. The device

has a slope efficiency of σs = 1 W/A and with _E = 1.5 eV follows from (6.15b)

a differential quantum efficiency of ηd = 0.66. The applied forward voltage for

diode lasers is

where Rd is the series resistance of the diode due to the resistivity of the epilayers

and the resistivity of the metal contacts. For typical devices the voltage drop

across the junction is 1.5 eV and one obtains V0 = 1.5V + i Rd.

The energy conversion efficiency of the diode laser, defined as the ratio of optical

output power to electrical input power, which we call the pump efficiency ηp,

follows from the above equations

The device characterized in Fig. 6.17 has an overall efficiency of ηp = 0.32 at

the maximum output based on a series resistance of Rd = 0.015_ and _E/e =

1.5 eV.

If quasi-cw bars are stacked into two-dimensional arrays, the bars are connected

electrically in series to avoid high currents and the associated resistive losses.

This is illustrated in Fig. 6.18. The drive current for this 10 bar stack is identical

to the current of a single bar. The nonlinear behavior of the overall efficiency

FIGURE 6.18. Performance of a quasi-cw 10 bar array (Spectra Diode Laboratories).

with increasing current occurs because the ohmic losses in the material increase

as the square of the current (see (6.15d)), whereas the output power is linear in

current.

Thermal Management

Diode-array pumping offers dramatic improvements in efficiency of solid-state

laser systems. However, the need to maintain the operating temperature within a

relatively narrow range requires a more elaborate thermal management system as

compared to flashlamp-pumped lasers.

In end-pumped systems, the diode wavelength is usually temperature tuned to

the peak absorption line of the laser and maintained at that wavelength by controlling

the array temperature with a thermoelectric cooler. This approach works

well for small lasers, and systems including thermoelectric coolers can work over

a large range of ambient temperatures.

In large systems, the power consumption of thermoelectric coolers is usually

prohibitive. For these systems, a liquid cooling loop with a refrigeration stage

can be employed that maintains the temperature of the coolant independent of the

environment. However, the coefficient of performance for refrigerators is typically

no greater than one. Therefore, the electrical input power requirement is equal to

the heat load to be controlled.

Recently, the development of diode arrays which can operate at high junction

temperatures has eliminated the need for refrigerated cooling loops for special

cases such as military systems. Diode arrays have been operated at junction temperatures

as high as 75◦ C, which is higher than ambient in most situations. A

simple liquid-to-air cooling system provides the most efficient thermal control

system for the laser because only the power consumption of a pump and possibly

a fan is added to the total electrical requirements.

The power dissipation of the various subsystems of a pulsed diode-pumped

laser is listed in Table 6.3. The system has a repetition rate of 200 Hz and an

output pulse energy of 0.8 J/pulse. By far the greatest heat dissipation occurs in

the diode pump arrays, and accounts for approximately 65% of the system input

power. The solid-state laser medium itself dissipates heat at a rate approximately

equal to the output power from the laser system. The electronics dissipate on the

order of 19% of the input power.

TABLE 6.3. Power dissipation from a

diode-array-pumped solid-state laser.

6.2.2 Side-Pumping with Diodes

In contrast to flashlamps, the emission from laser diodes is highly directional;

therefore many arrangements are possible for transferring the pump radiation to

the solid-state laser material. Since the output beams of laser diodes can be shaped

FIGURE 6.26. Flashlamp-pumped slab laser [18].

and focused, a major consideration is the design of a pump geometry which maximizes

the overlap of the pumped volume with the volume occupied by a low-order

resonator mode. The optimization of the overlap is referred to as mode-matching.

In side-pumping the diode arrays are placed along the length of the gain material

and pumping is transverse to the direction of the laser beam. Scalability to

higher power is the advantage of this design approach. As more power is required,

more diode arrays can be added along and around the laser rod. There are three

approaches to couple the radiation emitted by the diode lasers to the rod:

(a) Direct coupling. From a design standpoint this is by far the most desirable

approach. However, this arrangement does not allow for much flexibility in

shaping the pump radiation inside the laser rod.

(b) Intervening optics between source and absorber: In this case, the pump distribution

can be peaked at the center of the rod allowing for a better match

with the resonator modes. Optical coupling can be achieved by using imaging

optics such as lenses or elliptical and parabolic mirrors or by nonimaging

optics such as reflective or refractive flux concentrators.

(c) Fiber optic coupling. Due to the coupling losses combined with the increased

manufacturing cost of fiber-coupled diode arrays, this technique is not very

attractive for side-pumping.

Whether or not to use intervening optics between the arrays and the laser rod

depends mainly on the desired beam diameter and optical density of the active

medium. Generally speaking, oscillator pump heads usually employ optical elements

in order to concentrate the pump radiation in the center for efficient extractions

at the TEM00 mode, whereas large amplifier rods are direct-pumped.

FIGURE 6.27. Side-pumping of a cylindrical laser rod employing a hemispherical pump

geometry.

An example of a side-pumped Nd :YAG rode is shown in Fig. 6.27, the laser

rod is conductively cooled by a copper heat sink. The pumping cavity consists

of a number of linear arrays symmetrically located around the rod. The arrays

are mounted on long heat-sink structures and contain cylindrical optics. The optics

allow the arrays to stand-off from the laser crystal and the arrays. The output

apertures of the diode bars are imaged near the center of the rod by using a symmetrical

doublet of plano-convex cylindrical lenses. The rod surface in contact

with the heat sink has a reflective coating that returns the pump light for a second

pass through the rod.

For optimization of the pump distribution within the gain medium, a ray trace

program is very useful. Figure 6.28 presents the result of such a calculation for the

FIGURE 6.28. Contour plot of pump intensity distribution.

FIGURE 6.29. Arrangement of stacked diode array bars around a laser rod.

pump arrangement of Fig. 6.27. The code calculates the spatial pump distribution

in the laser material as a function of spectral absorption, and the spectral and spatial

properties of the radiation source as modified by the optical surfaces between

the source and absorber. The strong concentration of pump radiation at the center

of the rod was responsible for the good TEM00 performance of the laser.

A pump configuration of a high-power amplifier head which does not use coupling

optics is illustrated in Fig. 6.29. The laser rod is surrounded by eight stacked

arrays, each containing six bars. Since the bars are 1 cm long, and each bar produces

12 mJ in a 200μs pump pulse, the rod can be pumped at 576 mJ per centimeter

rod length.

Figure 6.30 displays a photograph of a pump head containing four rings of

stacked diode arrays. In this design, 4 cm of the 2 in long rod are pumped by the

diode arrays. Large diode-pumped lasers employ liquid-cooled laser rods or slabs

similar to flashlamp-pumped systems. The laser-diode arrays are directly mounted

on liquid-cooled heat sinks. Extremely compact and lightweight structures have

emerged which take full advantage of the high packaging density which can be

achieved in the diode arrays.

Figure 6.31 presents a photograph of an assembled pump head for a sidepumped

Nd :YAG rod pumped by 16 diode arrays. Each array is 1 cm long, and

the devices are arranged in four pairs symmetrically around the rod. The adjacent

four pairs are off-set 45◦ in order to achieve an eight-fold symmetry around

the laser crystal. The active pump area is 4 cm long and the complete assembly,

including the end plates, has a length of 7 cm.

Each diode array contains five bars, that is, the laser crystal is pumped by 80

bars at a pulse energy of 800 mJ in a 200μs-long pulse. Electrical input is 2 J per

pulse that at a repetition rate of 100 Hz amounts to about 180W of heat, which

has to be carried away by the laser head. To carry off heat from the diodes and

laser rod, liquid flows into the pump head through a fitting that screws into the

end plate. Directed by the manifolds, the coolant flows back and forth through

passages in the triangular structures, and also through the annulus between the

rod surface and sapphire sleeve. The flat areas of the triangular bars provide the

mounting surface for the diodes and the thermal interface. Eventually, the coolant

exits the pump head through a second fitting at the end plate.With the diode arrays

mounted on the internally water-cooled bars, the coolant, as it moves through the

structure, is close to the heat source, thus facilitating heat transfer. For maximum

FIGURE 6.32. Side-pumped conductively cooled slab laser.

heat transfer, a thin layer of indium is applied between the mounting surface and

the diode arrays.

Side-Pumped Slab Laser

Slab lasers are usually pumped by large densely packed two-dimensional arrays.

In a slab laser, the face of the crystal and the emitting surface of the laser diodes

are in close proximity and no optic is employed (Fig. 6.32).

The slab is usually pumped from one face only. The opposite face is bonded to

a copper heat sink containing a reflective coating to return unused pump radiation

back into the slab for a second pass. An antireflection coating on the pump face

is used to reduce coupling losses (the diode array is not in contact with the active

material). Liquid cooling is employed to remove heat from the crystal and diode

heat sink. The purpose of the zigzag optical path shown in Fig. 6.32 is to mitigate

the effects of thermal lensing caused by thermal gradients in the slab, as will be

explained in Chapter 7.

A photograph and details of a side-pumped and conductively cooled slab laser

are shown in Fig. 6.33. The zigzag slab Nd :YAG laser is air-cooled and batterypowered.

At a repetition rate of 10 pps, the system produces 270 mJ per pulse at

the frequency doubled wavelength of 532 nm. The slotted compartment shown in

the photograph houses the pump module consisting of the Nd :YAG slab, diode

arrays, heat sink, and cooling fans. The zigzag slab is pumped by twelve 18 bar

arrays with a total optical output of 2.6 J at a pulsewidth of 200μs.

A cross section of the pump module is shown in Fig. 6.33(b). The slab was

bonded to a thermal expansion-matched heat sink which is clamped to the bridge

assembly. This clamping method was employed to reduce mounting and thermal

cycling stresses in the laser crystal. Zigzag slab lasers can be scaled to the kilowatt

level employing diode arrays. The large face area available in a slab geometry

allows stacking of a sufficient number of diode arrays to achieve a high-output

power. The large flat surfaces eliminate the need for intervening optics between

the pump source and the laser crystal. This facilitates scaling to large systems.

The pump module of a high-power diode laser is shown in Fig. 6.34. A 20 ×

2 × 0.7 cm Nd :YAG slab is pumped by 14 diode arrays on each side. Each array

is comprised of 16 bars with a peak output of 50Wat a 30% duty cycle. The total

optical pump power to the slab is 6.72kW at an electrical input of 15kW. The

FIGURE 6.33. (a) Photograph of an air-cooled and battery-operated zigzag slab Nd :YAG

laser and (b) cross section of pump module [19].

system was designed for a pulse repetition rate of 2 kHz and a pump pulse length

of 150 μs.With a 2m long resonator, the pump head generated 740Wof average

Q-switched power in a 7 × diffraction limited beam. Peak power was 1MW.

Depending on the way heat is extracted from the slab we can distinguish the

three design configurations illustrated in Fig. 6.35. If the slab is water-cooled,

pumping and heat extraction can occur through the broad faces of the rectangular

slab as shown in Fig. 6.35(a). Space and military applications require conductively

cooled slabs. One approach, shown in Fig. 6.35(b), is to conductively cool one

face of the slab and pump through the opposite side. More powerful lasers can be

FIGURE 6.34. High-power diode-pumped Nd :YAG slab laser; shown is (a) the arrangement

of the diode arrays and (b) a cross section of the pump head [20].

built by conduction cooling of both large surface areas of the slab. Pump radiation

is introduced through the two narrow sides of the slab as shown in Fig. 6.35(c).

Therefore this pumping geometry is termed edge-pumping [21]. The concentration

of pump radiation, which is possible with diode sources, makes this approach

feasible. In all three cases depicted in Fig. 6.35 pump radiation enters into the slab

transverse to the resonator axis and the laser output is extracted from the ends of

the slab. Also the plane of the zigzag beam is collinear with the direction of the

thermal gradients in the slab.

Pump Configurations Based on Internal Reflections of the Laser Beam

Besides the classical side-pumped configurations discussed in this section, there

is an almost unlimited number of pump-configuration variations possible. Most

of the more unusual design configurations utilize internal reflections of the laser

beam. In Fig. 6.36 designs are illustrated which contain one, or many internal

reflections of the laser beam. An extension of this concept is the monolithic ring

laser which was discussed in Section 5.2.3.

FIGURE 6.35. Cross section of (a) liquid-cooled slab, (b) one-sided conductively cooled,

and (c) two-sided conductively cooled slab.

FIGURE 6.36. Pump configurations based on internal reflections of the laser beam. (a) Slab

with grazing angle at the pump face [22], and (b) slab with folded zigzag path [23].

FIGURE 6.37. Energy flow in a typical diode-pumped Nd :YAG laser. (Efficiencies are

defined in Section 3.4.1.)

In the design shown in Fig. 6.36(a) side-pumping through a polished face of

the gain medium is used to obtain high gain. The resonator mode makes a single

grazing-incidence total internal reflection at the pump face, thus remaining in the

region of the highest gain throughout its passage through the slab. This design

has been used for small Nd :YVO4 and Er : YAG lasers. Figure 6.36(b) exhibits a

miniature slab laser with a folded zigzag path of the laser beam. The entrance and

exit faces are cut at Brewster’s angle. The beam path in this particular Nd :YAG

slab consisted of five bounces. The pump radiation is collimated in the plane of

the zigzag path by a fiber lens to maximize the overlap between pump beam and

resonator mode.

The excellent spectral match of the diode output and the absorption of the solidstate

material, in conjunction with a high spatial overlap between pump radiation

and resonator mode volume, are responsible for the high overall efficiencies that

can be achieved in diode-pumped lasers. Figure 6.37 illustrates the energy flow in

a typical Nd :YAG laser, pumped with a 20W diode bar at 808 nm. Laser diodes

have efficiencies of ηp = 0.3–0.45, with quasi-cw diodes at the upper range

since they are more efficient pump sources as compared to cw diodes. Transfer of

diode radiation and absorption by the lasing medium is expressed by the transfer

and absorption efficiencies that in most systems are around ηtηa = 0.7–0.8. In

end-pumped configurations, diode radiation is, in most cases, not all collected by

the optical system and reflection losses further diminish pump radiation. On the

other hand, the diode radiation is almost completely absorbed in the longitudinally

pumped active medium. In side-pumped configurations, transfer efficiency

is close to one, but absorption of the pump radiation is often incomplete because

a portion of the diode output is transmitted through a small rod or thin slab.

The absorbed pump radiation produces heat in the crystal. The fractional thermal

loading, that is, the ratio of heat generated to absorbed power, is 32% in

Nd :YAG pumped at 808 nm. These losses are accounted for by the Stokes shift

and quantum efficiency ηsηQ = 0.68. Mode-matching and optical losses in the

resonator are expressed by the beam overlap and coupling efficiency which typically

are ηBηc = 0.6–0.8. The generic laser postulated in Fig. 6.37 has an overall

efficiency of 11.6% and an optical-to-optical efficiency of 35%. In reviewing the

energy-flow diagram, it is clear that the laser designer has only control over Steps

2 and 4, namely the design of the pump radiation transfer, pump geometry, and

mode volume.

6.2.3 End-Pumped Lasers

In the end-pumped technique, which is unique to laser diode pump sources, the

pump beam is circularized, and focused into the gain medium coaxial with the

resonator beam. There are two requirements that must be met for an efficient

TEM00 end-pumped laser. First, the gain medium must be long enough to absorb

all, or at least a large fraction, of the pump radiation. Second, the pump beam

spot size must be less than or equal to the spot size of the fundamental mode. If

these requirements are met and if we further assume that the pump beam and laser

resonator mode are TEM00 Gaussian beams, and the resonator length is equal to

the gain medium, a closed-form expression for laser threshold and slope efficiency

can be derived for a four-level system [24]. From the rate equations we obtain the

upper laser level population n at threshold

where Wp is the pump rate and rp(r, z) is the normalized pump intensity distribution,

τ f is the upper state lifetime, and n(r, z)/τf describes the losses due to

spontaneous emission. The pump rate is defined by

where Pa is the absorbed pump power, hνp is the pump photon energy, and ηQ is

the quantum efficiency.

From the threshold condition (3.7) and with (3.12) we obtain

the term s0(r, z) has been added to describe the normalized resonator mode intensity

distribution. From (6.23)–(6.25) follows

The spatial distribution of the pump volume rp(r, z) and resonator mode s0(r, z) is

taken into account by calculating an overlap integral that averages the upper laser

level population over the laser resonator mode. With the assumption of TEM00

Gaussian beams, the normalized expressions for the pump and resonator modes

are given by

α0 is the absorption coefficient at the pump wavelength, wp,w0 is the Gaussian

beam radius of the pump beam and resonator modes, respectively, and L is the

length of the gain medium and resonator. With (6.27) the overlap integral can be

calculated

With (6.28) following from (6.26) for the absorbed pump power at threshold

where hνp has been replaced by the laser photon energy hνL, and ηs is the quantum

defect efficiency.

From (6.29) it follows that a low-laser threshold requires small spot sizes and

low losses. The smallest mode dimensions are determined by diffraction and by

the length l of the resonator. As explained in Chapter 5, the cross-sectional area of

a mode is proportional to resonator length. Efficient absorption of pump radiation

requires that l should be on the order of or greater than 1/α0, where α0 is the

absorption coefficient for the pump radiation. With l determined by the pump

absorption length, the smallest values of wm and wp that can be achieved are

based on the propagation of Gaussian beams.

From (6.29) it follows that for a given laser mode volume the lowest threshold

is achieved for the smallest pump volume. The spot size wp(z) of the pump beam

inside the gain medium follows from (5.5):

where n0 is the refractive index of the gain medium and the beam waist w0p is

assumed to be at the exit face z = l of the laser crystal. The smallest pump volume

is achieved if the mean value w2p

is minimal over the crystal length. Integration

of (6.30) yields the smallest pump radius, which is theoretically possible for the

crystal of length

The power density in the resonator above threshold is obtained from the rate equations.

From this analysis the slope efficiency, defined as the ratio of laser output

to optical pump power for the case of wp < w0, is given by [24]

As in all lasers, low losses lead to a high slope efficiency. The above expression

is the same as (3.64) for the case where the pump radiation is completely

absorbed by the laser medium (ηa = 1), and perfect overlap between the inversion

and the resonator mode volume is assumed (ηB = 1). Since (6.32) gives

the optical to optical efficiency, the pump efficiency ηP and the throughput of

the optical transport system ηT, which occur in (3.64), are not included in this

expression.

The practical realization of a complete spatial overlap between the longitudinal

pump beam and the laser mode depends very much on the degree with which

the strongly astigmatic and highly divergent radiation emanating from the pump

source can be reshaped into a beam with circular symmetry and a small beam

waist. Focusing the output from a 1 cm bar or a stack of bars into a laser crystal

requires collimation of the emitted radiation, compensation of asigmatism and,

finally, focusing the beam to a circular spot.

Common techniques of coupling radiation from laser diode bars into the gain

medium of an end-pumped laser typically involve the use of:

• Two orthogonally mounted micro lens arrays for collimation of the output from

each emitter followed by a focusing lens (Fig. 6.38(a)).

FIGURE 6.38. Focusing of pump radiation into an end-pumped laser: (a) with two cylindrical

micro lens arrays and a focusing lens, (b) a single micro lens array and a focusing lens,

(c) a micro lens array followed by a fiber bundle and focusing optics, and (d) a nonimaging

light concentrator [25].

• A single cylindrical micro lens array for collimation of the fast axes of the bars

and a macro lens for focusing (Fig. 6.38(b)).

• A single cylindrical micro lens array for collimation of the fast axes of the

bars and a fiber bundle to circularize the beam followed by focusing optics

(Fig. 6.38(c)).

• A nonimaging light concentrator (Fig. 6.38(d).

• A macro lens system (Fig. 6.39).

A single laser bar or a stack of bars represent a pump source with an overall

cross section of up to a few centimenters squared, which consists of a multitude

of independent emitters. The smallest focused spot size from a single bar or

two-dimensional bar can be achieved if the radiation from each emitter is individually

collimated and corrected for astigmatism. This is usually done with micro

lenses or fiber bundles. After collimation, the multitude of beamlets are focused

with a bulk lens or lens assembly into the gain medium. This approach requires

extreme precision in positioning and aligning the optical components with the

emitters.

Designs based on ordinary macro lens assemblies or nonimaging light concentrators

treat the radiation emerging from a stack of bars as coming from a large

extended source. The designs are relatively simple but the minimum spot size

which can be achieved is at least an order of magnitude larger than can be obtained

by collimating the output from each emitter.

A compromise between the above mentioned approaches are designs whereby

the fast axis of each bar is collimated with a cylindrical micro lens and the emerging

beams are focused with a conventional lens or nonimaging light concentrator.

FIGURE 6.39. Optical system for end-pumping an Nd :YAG crystal with a stack of three

1 cm laser diode bars. Ray tracing plot (a) parallel and (b) perpendicular to the bars [27].

Lens Transfer System

A very effective method of collimating the output from two-dimensional bars

is by means of two orthogonally oriented arrays of micro lenses. The first set

of lenses are typically cylindrical micro lenses similar to short pieces of optical

fibers. Each lens is oriented along the length of a bar close to the facets of the

emitters. These fiber lenses, mounted in a rigid holder assembly, collimate the fast

axis of each bar within the stacked array. The number of micro lenses required is

thus given by the number of bars in the array. These fiber lenses are 1 cm long and

have a maximum diameter determined by the pitch of the array (typically 0.4 to

1 mm). The distance from the facets is determined by the back focal length which

is, for a micro lens with a circular cross section and uniform refractive index,

D(2 − n0)/4(n0 − 1). For a fiber lens with a diameter D = 800μm, made from

fused silica with n = 1.51, the distance between the micro lens and the facet is

therefore 200μm.

Instead of simple step-index fibers with a circular cross section, lenses with

noncircular cross section (aspherics) or graded index (GRIN) lenses can be used

to reduce spherical aberration. Aspheric lenses can be fabricated in large volume

from drawing of a shaped-glass preform. It is also common practice to make micro

lens arrays monolithic rather than from discrete segments of optical fibers.

Monolithic arrays consist of a glass plate in which parallel curved surfaces that

optically act the same way as individual fiber lenses are shaped (see Fig. 6.13(b)).

Mounted orthogonally in front of the first array of micro lenses is another set

of micro lenses for collimation of the slow axis or each emitter. The number of

micro lenses is given in this case by the number of emitters in each bar. The

diameter of the micro lenses cannot exceed the spacing between adjacent emitters

(0.2 to 1 mm). Because of the stand-off distance from the facets, which is largely

determined by the intervening horizontal lens array, the back focal length of the

second lens array must be significantly larger than the lens diameter. This rules

out micro lenses with a circular cross section since these optical elements have

a very short back focal length. Either aspheric or plano-convex cylindrical micro

lenses are normally used for the collimation of the slow axis of the emitters.

As an example, we consider a typical 1 cm cw bar which contains 20 emitters

each 100μm wide and spaced 400μm apart as shown in Fig. 6.12. If we assume a

5-bar internally cooled stack with a pitch of 1 mm, a collimator could be designed

that consists of a matrix of five horizontal cylindrical micro lenses with a circular

diameter of 800μm and a back focal length of 200μm and 20 vertical planoconvex

lenses with a diameter of 500μm and a back focal length of 2.5 mm.With

the approach discussed so far, the output from a stacked bar can be focused to a

spot with a diameter of a few hundred micrometers.

If larger spot sizes can be tolerated, a compromise between complexity and

performance can be reached if only one array of micro lenses is employed followed

by an ordinary lens system. In a typical design, a set of fiber lenses reduces

the beam divergence of the fast axis of each bar to match the divergence of the

slow axis and a bulk lens focuses the beam into the active medium. Figure 6.38(b)

illustrates such an approach.

In the absence of micro lenses or fiber-optic bundles one can attempt to focus

the radiation from a single bar or stacked bar array with conventional optics. As

already mentioned, the power density at the focal point will be far below that

which can be achieved with micro-optics. To illustrate this point and to provide a

comparison the following example is given.

Figure 6.39 depicts an imaging optical system for focusing pump radiation from

a stack of three 1 cm bars into an Nd :YAG crystal. The optic consists of two

cylindrical lenses, L1 and L2, with 12.7 and 19mm focal length, respectively,

followed by a 30mm focal length plano-convex lens L3 and a 12mm focal length

aspheric L4. The rays traced in the plane perpendicular to the bars span a total

angular field of 60◦ at the source and those traced in the plane parallel to the bars

cover a total field of 10◦. The corresponding width at the pump end of the laser

rod is ≈ 2 mm.

In order to at least approximately match such a large pump profile with the

resonator mode, a near hemispherical resonator was chosen. The resonator length

was adjusted until the TEM00 mode at the curved mirror and the front surface

of the Nd :YAG was expanded to 1.8 mm. At the flat output coupler the 1/e2

beam diameter was 120μm. The expansion of the mode at the pump end of the

resonator is achieved at the expense of alignment stability. A photograph of the

system is shown in Fig. 3.27.

Fiber Coupling

Using fiber optics to deliver the diode output to the laser crystal has a number of

advantages: the beam at the fiber output has a circular distribution and the ability

to remove heat from the diode remote from the optical components of the laser

is an attractive feature. Also the pump source can be replaced relatively easily

without disturbing the alignment of the laser. Typically one of the following three

techniques is implemented in coupling the output from a diode bar or stacked bar

into an optical fiber.

The first approach is identical to the technique already described for endpumping

a laser. Two sets of micro lenses and a condensor lens are used for

focusing the radiation into a multimode fiber with a diameter of 200 to 400μm.

In another design, the output from a diode bar is first collimated using a fiber

lens and subsequently imaged with a combination of lenses into a two-mirror

beam-shaping device. With this focusing scheme the output from a 1 cm bar has

been focused into a circular spot of less than 200μm [26].

In the third approach a fiber bundle is employed to transform a line source,

which a single 1 cm bar represents, into a circular beam. A micro lens reduces the

divergence of the fast axis of the emitters before the radiation is collected by a

flat bundle of fibers (see Fig. 6.38(c)). The line-to-bundle converting fiber-optic

coupler consists of a large number of small fibers spread out in a linear array

which faces the emitting diodes of the laser bar. At the other end, the fibers are

arranged into a round fiber bundle.

The technique of using a fiber-optic bundle to circularize the output from a

1 cm diode bar is implemented in the laser illustrated in Fig. 6.40. Shown is an

Nd : YV04 laser crystal pumped from both ends by two fiber-coupled 20W diode

bars. For separation of the pump and resonator beam, two dichroic beam splitters

are inserted into the optical train, which results in a z-configuration of the optical

resonator. The radiation from each bar is coupled into a flat fiber bundle by a

micro lens, which reduces the divergence in the fast axis of the diode output. At

FIGURE 6.40. (a) Optical schematic of an Nd :YVO4 laser pumped with fiber-coupled

diode bars (Spectra-Physics Lasers, Inc.) and (b) laser output versus optical pump power

incident on the crystal.

FIGURE 6.41. End-pumping with a nonimaging light concentrator [28].

the pump end, the fibers form a round bundle, and a collimating and focusing

lens transfer the pump radiation into the crystal through a dichroic fold mirror.

The high efficiency that can be achieved with diode end-pumped solid-state lasers

is illustrated in Fig. 6.40(b). With a fiber coupling efficiency of 85% and some

reduction in the rated power level of the diode to preserve lifetime, each side

of the Nd :YVO4 crystal is pumped with 13W. Pump radiation at the crystal is

converted with over 50% conversion efficiency into a polarized TEM00 mode.

Maximum output of the system is 13W. An Nd :YAG laser used in the same

configuration produced 11W of polarized cw output.

Nonimaging Light Concentrators

Figure 6.41 illustrates an end-pumping technique based on a nonimaging devices.

In this approach the output from a stack of bars is first collimated in the fast axis

by cylindrical micro lenses mounted in front of each bar. The collimated radiation

is directed into a lens duct that combines focusing at its curved input face and total

internal reflection at the planar faces to channel the diode pump radiation to the

laser rod. Powerful Nd :YAG, Tm:YAG, and Yb :YAG lasers have been pumped

with this technique and outputs in excess of 100W have been reached in these

lasers. A slightly different version of a light concentrator that has parallel walls

in the vertical direction is shown in Fig. 6.38(c). Concentrators are an attractive

solution for large stacks because they are relatively easy to fabricate. The degree

of concentration is governed by the radiance theorem stating that the radiance of

the light distribution produced by an imaging system cannot be greater than the

original source radiance.

6.2.4 Face-Pumped Disks

In a disk laser amplifier the surfaces are set at Brewster’s angle, as shown in

Fig. 6.42. The slab faces have a high-quality optical finish to minimize scattering

loss. The beam is linearly polarized in the p-plane to avoid reflection loss.

Flashlamps are used for pumping, with lamp radiation incident on the disk faces

and transverse to the beam direction. Thus, the disks are said to be face-pumped.

Nd : glass amplifiers containing glass disks have mainly been developed at labo

FIGURE 6.42. Schematic diagram of (a) a disk amplifier and (b) alternate arrangement of

disks.

ratories engaged in laser fusion studies. Disk amplifiers with apertures as large as

74 cm have been built and operated.

Generally, an even number of disks is used in an amplifier since beam translation

occurs due to refraction in a single disk. Beam steering can be minimized if

the disks are placed in an alternate rather than parallel arrangement, as shown in

Fig. 6.42(b).

The concept of face-pumping, as opposed to pumping through the edges of a

disk or slab, is important in that the pump radiation can be distributed, in an easily

controlled manner, uniformly over the faces of the slab. Pumping in this way

yields uniform gain over the aperture for a laser beam, but more important, the

slab heating is uniform laterally, thus allowing transverse temperature gradients

to be minimized. Such temperature gradients give rise to distortion in the slab just

as in the rod geometry.

The design of the flashlamp reflector housing requires a careful design to provide

uniform illumination of the slabs. The cusp-shaped reflectors direct radiation

to the slabs and shield each flashlamp from the radiation from adjacent lamps.

Table 6.5 lists the results of an energy-transfer analysis performed on a large

Nd : glass disk amplifier. About 1% of the total electrical input is available as

stored energy at the laser level. Of course, only a fraction of this energy can be

extracted during the amplification of a laser pulse.

The other extreme of a face-pumped laser, in terms of size, is the thin-disk

approach employed for pumping Yb :YAG. As illustrated in Fig. 6.43 a disk about

TABLE 6.5. Energy transfer in an

Nd : glass disk amplifier [29].

300μm thick is pumped in a small spot of a few millimeters in diameter at a small

angle about the normal of the disk face. The 941 nm pumps radiation reflected

back from the disk is refocused by a mirror for a second path through the crystal.

The backside of the disk, which is in contact with a heat sink, is coated to form one

of the resonator mirrors. The other resonator mirror is the curved output coupler

shown in Fig. 6.43.

In the thin-disk laser, waste heat generated in the disk is extracted through the

back high-reflectivity mirror and to an actively cooled submount that holds the

laser disk. The key feature of the thin-disk design approach is that this cooling

geometry constrains resulting thermal gradients to lie essentially parallel to the

laser resonator axis. Consequently, thermal focusing by the gain disk is greatly

reduced, facilitating the achievement of high beam quality.

FIGURE 6.43. Thin-disk-pumping geometry of a Yb :YAG laser [30].

Summary

This chapter deals with the design and characteristics of the pump sources employed

in solid state lasers and with the techniques and methods that have been developed

for transferring the radiation from the pump source to the laser medium.

The designs of an optical pump system for a solid state laser is governed by

the following considerations: the spectral and spatial emission characteristics

of the pump source, transfer of pump radiation to the gain medium, absorption

of the pump radiation, and spatial overlap with the mode volume in the

resonator.

The pump sources for solid-state lasers are flash lamps, continuous arc lamps,

and laser diode arrays. In flash lamps that are typically operated at high current

densities, the spectral line structure of the particular fill gas is masked by a strong

continuum. The output spectrum resembles a blackbody radiator. In continuous

arc lamps, the continuum emission spectrum contains sharp emission lines that

are characteristic of the particular fill gas. The output from flash lamps and arc

lamps is broadband, and only a small fraction of the total output spectrum falls

within the absorption bands of the active medium. Diode lasers, on the other hand,

are narrowband emitters and by adjusting the composition of the semiconductor

material, the output wavelength can be shifted to match an absorption band of a

particular solid-state laser material.

The transfer of radiation from the pump source to the laser medium should

be highly efficient, and ideally the pump radiation should create an inversion

that spatially overlaps the volume of the resonator mode. The total absorption

of pump radiation depends on the absorption coefficient and thickness of the laser

medium. However, mode structure and thereby beam uniformity and divergence

are strongly affected by the pump power distribution inside the active material.

In flash lamps and arc-lamp pumped lasers, the active medium and the radiation

source are contained in highly reflective enclosures. Typical pump cavities consist

of reflective elliptical cylinders with the lamp and active medium at each focal

axis, or so-called closed coupled cavities where the lamp and laser rod are in close

proximity surrounded by reflective metal or white ceramic walls.

Since radiation from diode lasers is directional, reflective enclosures are not

needed. In the so-called side-pumped configuration, the laser diode arrays are

placed along the length of cylindrical laser rods or rectangular slabs and the pump

radiation propagates perpendicular to the axis of the resonator. In this configuration

laser diode arrays are essentially used as replacements for flashlamps.

Other designs exploit more fully the spatial properties of laser diodes by employing

intervening optics to shape the pump beam to achieve good overlap with the

resonator fundamental mode. The most efficient pump method is end-pumping,

whereby pump and laser radiation propagate collinear inside the active material.

End pumping has the advantage that total absorption of the pump radiation is obtained,

and an excellent overlap between the pump volume and the laser mode can

be achieved.